## 사양서EEG-1200K
COMMODITY DESCRIPTION
關 稅 分 類 番 號 |  |  |  | 品 目 및 規 格 | 單 位 | 數 量
H.S.NO. |  |  |  | Description | Unit | Quantity
9018.19.1000 |  |  |  | Electroencephalograph | SET | 1
EEG-1200K
A. 특징
| 1. | DSA Trendgraph
뇌파 파형의 주파수 성분을 분석하여, 각 주파수에 해당하는 파형진폭을 DSA(Density Spectral Array)
트랜드 그래프를 이용하여 도시화 한다. 이 그래프를 통해 장시간의 뇌파파형에서 이상징후가 나타난
특정 부분을 손쉽게 찾아낼 수 있다.
| 2. | 64 channel display : 64채널의 EEG 파형을 디스플레이 한다.
| 3. | USB Interface
뇌파junction box 와 컴퓨터 본체의 통신방법으로 USB 통신 환경을 채택하여, 신속 정확한 신호전달체계를
유지한다.
| 4. | Network capability
기본적인 네트웍 작업환경이 가능하여 측정실과 판독실 및 여타의 네트웍 환경으로 데이터 공유가 가능하다.
| 5. | Large data storage
고용량의 저장공간을 확보하고 있으며, CD 디스크에 영구적으로 데이터를 보관할 수 있다.
| 6. | Automatic photic stimulation
자동 프로그래밍 광자극 및 수동 광자극 기능이 가능하다.
| 7. | Data review during acqusition
뇌파 측정과 동시에 파형 리뷰가 가능하여 측정과 편집을 효율적으로 시행할 수 있다.
| 8. | File append : 분리된 여러 파형을 하나의 파형으로 편집할 수 있다.
| 9. | Montage Map : 전극 위치가 표시된 몽타쥬 맵이 화면상에 나타난다.
| 10. | Note window
드래그 앤 드롭 방식으로 특정 파형을 복사하여 비교 분석할 수 있다.
| 11. | EEG PortaView
EEG PortaView 프로그램을 이용하여 데이터를 다른 컴퓨터에서 특별한 소프트웨어 없이 리뷰할 수 있다.
| 12. | 3D Voltage mapping
6가지 다른 양상의 voltage map을 볼수 있다.
B. 사양
| 1. | 데이터 획득
|  |  | (1) | 입력 채널 수
- 전극위치에 따른 뇌파 입력 : 25
- 추가 입력 : 4(Bipolar : X1 to X4), 8(Monopolar : X1 to X4)
- 피드백 기준 입력 :1
- 호흡측정 입력 : 3
- 양극 입력 : 6(3pairs)
- DC 입력 : 4
- SpO2 입력 : 1
- CO2 입력 : 1
|  |  | (2) | 입력 임피던스 : 100㏁
|  |  | (3) | 입력 누설전류 : 5nA 이하
|  |  | (4) | CMRR : 105dB or greater
|  |  | (5) | Low cut filter : 0.08Hz
|  |  | (6) | High cut filter : 120Hz
|  |  | (7) | A/D Conversion : 16bits
|  |  | (8) | 샘플링 주파수 : 100,200,500,1,000Hz
| 2. | 데이터 처리
|  |  | (1) | 감도
- EEG 입력 : Off, 1,2,3,5,7,10,15,20,30,50,75,100,150,200㎶/mm
- DC 입력 : Off, 10,15,20,30,50,70,100,150,200 mV/mm
|  |  | (2) | 시간 상수 : 0.001,0.003,0.03,0.1,0.3,0.6,1.0,2.0s
|  |  | (3) | Low cut filter : 0.08,0.16,0.27,0.53,1.6,5.3,53,159Hz
|  |  | (4) | High cut filter : 15,30,35,67,70,120Hz
|  |  | (6) | AC interference filter : 50 or 60 Hz
|  |  | (7) | 조정파형
Waveform shape : 0.25Hz step wave or 10 Hz sinewave
Voltage : 2, 5, 10, 20, 50, 100, 200, 500, 1000uV(x1000 for DC input signal)
|  |  | (8) | 임피던스 점검
|  |  | (9) | Pattern : 36 sets of programmable montages combined with programmable
individual amplifier settings
|  |  | (10) | 기준전극선택 : A1->A2, A1<-A2, A1<->A2, A1+A2, VX,
AV(with unsuitable electrode deletion function), Aav, Org, SD, and OFF
|  |  | (11) | 마킹 신호 : 광자극 마크, 과호흡 마크
| 3. | 화면표시
|  |  | (1) | 해상도 : 1024×768 pixel
|  |  | (2) | 디스플레이 채널수 : 64 채널
|  |  | (3) | 디스플레이 칼라 : 16 colors
|  |  | (4) | 파형위치 조정 기능
|  |  | (5) | 파형 일시정지 기능
|  |  | (6) | 파형 속도 : 5,10,20,30,60 s/page or 5min/page
|  |  | (7) | 시간 마크 : 0.1,1s
|  |  | (8) | 이벤트 마크 : 화면 하단 표시
|  |  | (9) | EEG 스케일 조정 기능
| 4. | 측정모드
|  |  | (1) | Time mode : Manual timer, Recording timer, HV timer
|  |  | (2) | Timer counting rate : Up to 99 min 59s
|  |  | (3) | Sampling frequency : 100, 200,500,1000Hz
| 5. | 리뷰모드
|  |  | (1) | 조정가능 아이템 : 몽타주, 감도, 시간상수, high-cut filter
기준전극, 표시속도
|  |  | (2) | 바로가기 기능 : 특정 이벤트, 페이지, 특정시간
|  |  | (3) | Display mode : Continuous, high speed, high speed with pause
manually page by page, manually second by second, and manually waveform centering
|  |  | (4) | Display information : Event, Channel number, montage, and comment
| 6. | 광 자극
|  |  | (1) | Lighting intensity : 4.0lx
|  |  | (2) | Stimulation modes : 3 automatic, manual, and single
|  |  | (3) | Mode Operation : Continuous operation with intermittent loading
|  |  | (4) | Duty cycle : Max. 5 minutes continuous operation in 30 min
|  |  | (5) | Automatic stimulation
- Stimulation rate : 0.5,1 to 33, 50 and 60Hz
- Stimulation period : 1 to 99 seconds
- Pause period : 1 to 30 seconds
|  |  | (6) | Manual stimulation : Manually set frequency and stimulatiro period
- Photic frequency : 0.5, 1 to 33,50 and 60Hz
- Stimulation tome : 1 to 99s and continuous stimulation
- Pulse mode : Normal, random and double
- Random stimulation : 1 to 33 Hz
- Single stimulation : Manual and automatic automation by external signal
| 7. | 안정성
|  |  | (1) | Safety standards : IEC60601-1(1988), IEC60601-1Amendment 1(1991),
IEC60601-1 Amendment 2(1995), IEC60601-2-26(1994), EN 60601-1-1(1992-06)
with AM1(1995-10)
|  |  | (2) | Type of protection against electric shock : Class I
|  |  | (3) | Degree of protection against electric shock : Type BF
| 8. | 전자기적 호환성
|  |  | (1) | IEC60601-1-2 (1993)
|  |  | (2) | CISPR11 Group 1 Class B
| 9. | 전원사양
|  |  | (1) | Line voltage, Line frequency : AC 220V to 240V, 50 or 60 Hz
|  |  | (2) | Power consumption : 750VA
C. 구성
|  |  | 1. | Main PC Unit | 1 SET
|  |  | 2. | Electrode Junction box | 1 SET
|  |  | 3. | System program | 1 SET
- Routine EEG Measuring s/w
- 3D Voltage Mapping
- Frequency Mapping
- 8ch DSA Trandgraph display
- EEG Scope function
- EEG Review program
- EEG PortaView program
|  |  | 5. | Flash lamp assembly | 1 SET
|  |  | 6. | Photo control unit | 1 SET
|  |  | 7. | Flash lamp stand | 2 SET
|  |  | 8. | Cart (Domestic Supply) | 1 SET
|  |  | 9. | Standard Accessory Set | 1 SET
- Collodion EEG disk electrode set
- EEG ear clip electrode set
- EEG paste, ELEFIX, 180g
- ECG electrode lead
- ECG clip-on electrode
- ECG paste, CardioCream
- Marker cord
- Power cord
- Ground lead
- Fuse
|  |  | 10. | Operating manual & Service manual(CD) | 1 SET
D. 비고 |  |  | 1. | 장비설치 및 사용법 교육은 공급자가 시행한다.
|  |  | 2. | 하자보증기간은 장비설치 및 정상작동 테스트 이후 2년으로 한다.